Push-in RTD temperature probes(TM1101)

Applications：

The Push-in RTD temperature probes are used to measure the temperature of gases or liquids, including HVAC, Refrigeration equipments, Heaters, Ovens, Mechanical equipments and Experimental facilities, etc.

Technical data:

Connecting cable:  PVC, temperature range:-5℃~+85℃

Silicone, temperature range:-50℃~+180℃

Teflon, temperature range -50℃~+260℃

Glass Fiber, temperature range:-50℃~+350℃

Protection tube:    Stainless steel 304/316, OD=4mm,5mm, 5.2mm,6mm

Sensor element:    Pt100, Pt500, Pt1000 meets IEC 751

Tolerance :        Class B: △t=±(0.30 + 0.005 |t| )℃

Class A: △t=±(0.15 + 0.002 |t| )℃

Response time：t0.9  approx. 10secs, in water 0.2 m/sec, diameter 6mm

接线图 Wiring diagram

铂电阻 RTD

单支2线 1x2-wire

单支3线 1x 3-wire

单支4线 1x4-wire

双支4线 2x4-wire

双支6线 2x6-wire

白色芯线 white core wire

红色芯线 red core wire

Dimensions(外形尺寸图）

Order Details:

| TM1101 | — | (1) | — | (2) | — | 000 | — | (4) | — | (5) | — | (6) | / | (7) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

- Protection tube diameter D

4     4mm

5     5mm

5.2   5.2mm (EL≤45mm）

6     6mm

- Insertion length EL

30    30mm

45    45mm

50    50mm

Note：other lengths as per customers’ requitements

- (4)  Sensor element

11     1 x Pt100 in 2-wire circuit

51     1 x Pt500 in 2-wire circuit

61     1 x Pt1000 in 2-wire circuit

13     1 x Pt100 in 3-wire circuit

14     1 x Pt100 in 4-wire circuit

12     2 x Pt100 in 4-wire circuit

16     2 x Pt100 in 6-wire circuit

62     2 x Pt1000 in 4-wire circuit

(5)  Measuring range/connecting cable

10    -5℃ ～ +85℃ /PVC

13    -50℃ ～ +180℃/Silicone

14    -50 ℃～ +260℃/Teflon

15    -50 ℃～ +350℃/Glass Fiber

- (6) Cable length  L（L≥100mm）

100    100mm

200    200mm

2000   2000mm

Note：other lengths as per customers’ requitements

(100mm≤L<1000mm ,L1=50mm；L≥1000mm, L1=100mm)

- (7) Additional options

00     No

01     Strain relief spring/ strain relief spring hose

02     Shielded connecting cable (D≥4mm)

03     Waterproof function

07     Ferrules as per DIN 46 228 part4 (Standard)

08     Multipole connector (Providing types)

99     Special options

List the additional options in sequence, separated by commas.